## Slide 1
GUI-RITS(ver1.5.0) profile shape memo 05/28/2025
Change slide p4 for ver1.5.1, 2025/08/28

## Slide 2
Convolution of the Jorgensen function with the Heaviside’s step function
Jorgensen type Bragg edge profile function
You can find these mathematical formula in the reference paper.
Notes: Aim of the data analysis code is the quantification of various crystalline structural information in Bragg-edge transmission imaging. (+0:1 min = 1:2 min)

## Slide 3
Jorgensen function for TOF diffractometry
Broadening parameters
width
rise
decay
You can also find these mathematical formula in the reference paper.
But you may not find these mathematical formula in the reference paper.
Notes: Aim of the data analysis code is the quantification of various crystalline structural information in Bragg-edge transmission imaging. (+0:1 min = 1:2 min)

## Slide 4
Facility | Moderator | Sigma0 | Sigma1 | Alpha | Beta
1 | Coupled | x10.0 | x10.0 | x0.01 | x0.001
2 | Decoupled | x10.0 | x10.0 | x0.01 | x0.01
3 | Poisoned | x1.0 | x1.0 | x0.01 | x0.01
4 | LARMOR | x10.0 | x10.0 | x0.001 | x0.001
This is the only one that is not a flag for refinement. Basically on.
Refine after the lattice constant close enough.
There is no need to be wastefully wide. About this much.
“Decoupled” has to be selected for BL10 and BL22 at J-PARC. If you make a mistake, the parameters will change by a factor of 10.
Broadening parameters
Neutron pulse width
Neutron pulse rise
Furthermore, in the case of RITS, σ is supposed to be 1/10 of its true value, and α and β are supposed to be 100 times their true values, which makes it a bit confusing.
Neutron pulse decay
Crystallite size
Instrument and micrstrain
enter the value of t0 if you need.

## Slide 5
In this curve-fit, sigma0 and sigma2 were fixed to 0.1 and 0, and the converged value of sigma1 was 0.24(9).
t = λ × 252.778 × L1+ t0
We fix t0 = 0.
In these curve-fit, alpha0, alpha1, beta0 and beta1 were converged to these values, respectively.
Individual peak fit of standard iron sample(s).